		Co-ordination of Notified Bodies Medical Devices / (NB-MED) / on Council Directives 90/385/EEC and 93/42/EEC	Recommendation NB-MED/2.5.1/Rec5
Title:	Technical Documentation

# Introduction

The requirements for the technical documentation are laid down in the various An- nexes of the medical devices Directives, as appropriate for the conformity assess- ment procedure and the products concerned. As a general rule, the documentation should cover the design, manufacture and intended operation of the product.

Note. The ‘operation’ of the product would include installation, preparation for use, pre-use checks and maintenance, calibration and servicing as appropriate to the particular medical devices involved.

The details included in the documentation will depend on the nature of the product and on what is considered as necessary, from a technical point of view, for demon- strating the conformity of the product to the essential requirements of the medical devices Directives. If the harmonised standards have been applied, the technical documentation should also make clear where these have been used to demonstrate conformity with the particular essential requirements covered by the standards.

Note: This Recommendation has been written particularly to meet the needs for guidance on technical documentation for medical devices and active implantable medical devices. The Recommendation may also be helpful, however, in relation to IVDs, but may need revision in the light of experience of the practical implementation of the IVDD.

# Purpose of Recommendation

The purpose of this Recommendation is to provide guidance to Notified Bodies, Competent Authorities and manufacturers on the technical documentation needed to meet the requirements of the medical devices Directives.

Note: It is not the purpose of this Recommendation to oblige the manufacturer to re- organise existing technical documentation where this already proves appropriate and sufficient.

# Technical Information

  1. General

VdTÜV
Technical Secretariat NB-MED PO Box 10 38 34
D-45038 Essen
Hinrich Schaub (- 178) Jörg Höppner (- 138) Kurfürstenstraße 56
D-45138 Essen
Phone:++49/201/8987- 0
Fax:	++49/201/8987- 120
eMail: vdtuev.hoeppner@t-online.de
vdtuev-document dn: ...\hoeppner\mp\nb\rec_vdt2\R2_5_1_5.doc_rev4.doc

Subject	Guidance

  1. Content of technical docu- mentation
This Recommendation is not intended as an ex- haustive listing of all the technical documentation that may be required in particular cases. Equally, where particular information indicated by this Recommendation is not included, a justification should be provided.

The manufacturer is required by the Directives to prepare the technical documentation and to de- termine and justify what is appropriate and suffi- cient to assure his particular medical devices comply with the relevant Directive(s). This will obviously vary on a case-by-case basis, de- pending on the type of product, the risk associ- ated with its manufacture, installation, use and servicing, and the period that it has been on the market. For example, it is unlikely that well- established products, regardless of their classifi- cation, will have much formal design validation - but the manufacturer will have considerable market experience of use, together with details on how he has responded to any problems that have emerged, and this available data should be used by way of validation.

  1. Other	The technical documentation should include the information indicated in sections 3.2 - 3.5

# Product Description

Subject	Guidance

a general description of the device(s)
This should include the intended range of vari- ants (for example, a group of catheters of a par- ticular type differing only in length), and a de- scription of the packaging where this is relevant to the preservation of the intended characteris- tics and performances of the device(s). All that is needed, however, is a brief description sufficient to allow an understanding of the design, char- acteristics, and where appropriate, performances of the device(s) and to distinguish between vari- ants. In many cases, the name of the device(s)

will be sufficient. Where the ‘technical docu- mentation’ is to be evaluated by a Notified Body, a general pictorial representation of the device(s), e.g. a schematic diagram, photograph or drawing may be of assistance.

Note: For guidance on what constitutes a ‘range’ of product variants, see NB-MED Recommenda- tion NB-MED/2.1/Rec2.

a description of the intended use and operation of the device(s)
A short description of the intended pur- pose/application and/or method of use of the device(s) is needed. This may include, where appropriate, details of the intended patient population(s) and medical condition(s) for which the device is intended. This should also make clear the intended user(s), in particular whether the device is professional use. All of the above may be self-evident from the general description of the device(s). The information may be given by way of reference to the ”instructions for use” or operating manual for the device(s). It is not necessary to detail the mechanism by which the device(s) achieves its intended purpose.

The description should make clear the process- ing of important inputs and outputs associated with the intended operation of the device(s). For example, in the case of a pulse oximeter, the inputs and processing are to do with relative ab- sorption of light at different wavelengths. The outputs may include a display of the haemoglo- bin saturation with oxygen, and a graphical representation of the patient’s pulse.

device(s) incorporating a medicinal substance
Where the device(s) incorporates a medicinal substance, the ‘technical documentation’ should make clear the purpose of including the sub- stance and its mode of action in this application. This only applies where the substance is liable to act upon the body with action ancillary to that of the device.

Note: For this reason this will not be relevant to IVDs which do not ‘act upon the body’.

The risk analysis should address the additional risks and benefits associated with incorporation of such a substance. The technical documenta- tion should include the data on the tests con- ducted in this connection.

Note. For guidance on ‘borderline products’ and regulatory treatment of ‘drug/device combina- tions’ see MedDev 2.1/3.

device(s) incorporating non- viable materials of animal ori- gin
Where the device(s) incorporates non-viable materials of animal origin, the risk analysis within the technical documentation should address the additional risks and benefits associated with in- corporation of such materials, and the measures taken (for example, in sourcing of animals, vet- erinary controls and measures taken to elimi- nate/inactivate transmissible agents).

Note: For guidance on measures to elimi- nate/inactivate transmissible agents see MedDev 2.5/5.

Note: IVDs may contain materials of animal [or human] origin. Here, the technical documenta- tion should include relevant details, including sourcing and measures to protect personnel and to preserve the performance of the device.

device(s) requiring special consideration
Where aspects of the device(s) are the subject of emerging concern (for example, the use of latex potentially leading to allergic reaction), the risk analysis within the ‘technical documentation’ should address these aspects.

description of the methods of manufacture envisaged
A summary is required in general terms of the type of manufacturing method (for example, injection/blow moulding, extrusion, chemical processing, assembly, packaging/labelling) and the method of sterilization, if relevant. This should make clear the technologies involved and means of assuring the intended characteristics and performances of the devices manufactured. What is not required is an exhaustive description

of manufacturing processes.

description of the accessories, adaptors and other devices or equipment and other interfaces which are intended by the manufacturer to be used in combination with the device(s)
The technical documentation should include the description of other devices or equipment etc. which the device is intended to be used with; for example, where the manufacturer makes specific claims concerning compatibility. It should also include data on the verification and validation of the safety and performance of such combina- tions.

In describing the requirements for safe and proper operation of the device(s) when used in combination with devices or equipment from other manufacturers, what is needed is a brief description sufficient to understand the important parameters or interfaces (for example, the con- nectors needed or the voltage, frequency and/or stability of the electricity supply required).

The technical documentation should also address known incompatibilities which may be covered, for example, in the label or the instruc- tions for use.

classification of the device under the relevant Directive
The technical documentation should include the rule number(s) applied under the Directive, together with a brief rationale for this classifica- tion, and reasons why particular rules do not apply, if this is not self-evident.

In the case of IVDs, the classification of a par- ticular device is self-evident from the lists given in Annex II of the IVDD or where the device is labelled as for ”self-testing”.

# Technical Requirements

Subject	Guidance
Identification of technical requirements
The manufacturer should make clear the Directive(s) which apply to the particular device(s) concerned, including Directives other than the medical devices Directives. Where not

self-evident, the manufacturer should document the rationale for classifying as a medical device and deciding what other, if any, requirements apply.

In each case, those essential requirements (ERs) of the Directive(s) (e.g. MDD Annex I, AIMD Annex 1, IVDD Annex I) and other requirements of the Directives which apply should be identified. The manufacturer should define the technical requirements/specifications which must be satisfied in order to ensure that each of the applicable Directive requirements are met.

Where particular ERs are deemed not to apply to the device(s) concerned, a brief rationale should be given where this is not self-evident.

Solutions adopted to fulfil the essential requirements
The manufacturer is required by the Directives to demonstrate how each of the applicable essen- tial requirements and any derived technical requirements/specifications for the particular device(s) concerned has been met.

Compliance with published standards is volun- tary.

Where ”harmonized standards” are used to comply with relevant essential requirements, all that is needed is to demonstrate the device(s) concerned complies with the relevant clauses of the ”harmonized standard(s)”.

Where other methods, including compliance with draft and in-house/industry standards, are used to comply with one of a range of relevant essen- tial requirements, the manufacturer should justify that:
1. the methods applied adequately address relevant essential requirement(s) and
2. the device(s) concerned comply with the relevant provisions of these.

The evidence of device compliance with stan-

dards may take the form of, for example, test re- ports or records of application of Standard Operating Procedures (SOPs) intended to assure such compliance.

Note: The use of a checklist may facilitate demonstration of how the solutions adopted meet the relevant requirements (”Essential Requirements Checklist”). Such a checklist should:
  1. list the essential requirements, identifying those which are/are not applicable
  2. list the standards applied, and
  3. against each essential requirement, give the basis for claiming compliance. This will either directly make clear the solutions adopted to fulfil each requirement or refer to stand-alone specifications, reports and the like.

Standards applied	Where the manufacturer demonstrates confor-
mity with particular essential requirements by claiming compliance with available published standards, the Directives require that these standards should be identified. The manufac- turer should make clear where standards which are applied in full or in part are ”harmonized standards” (including ”common technical specifi- cations” in the case of IVDs). Compliance with all or parts of such ”harmonized standards” carries the presumption of conformity with relevant essential requirement(s) of the Directive(s).

Where device(s) do not comply with key relevant published standards, a rationale should be given.

Note: Draft standards may also be used as guid- ance, but here the manufacturer should have regard to how these may change prior to publi- cation.

Note: Certain Pharmacopoeial Monographs have a status equivalent to that of ”harmonized stan- dards”.

# Design

Subject	Guidance

the results of the risk analysis	The manufacturer is required by the Directives to
present the documented results of the risk analysis.

The risk analysis should address all hazards known or reasonably foreseeable for the par- ticular product types and technologies involved, together with the likelihood and consequences of occurrence and measures taken to reduce the resulting risks to acceptable levels. This should address all relevant risks. For example, in the case of devices incorporating e.g. a medicinal substance or materials of animal origin, or natu- ral rubber latex, the risk analysis should include the additional risks and benefits associated with incorporation of such substances.

In the case of devices intended and labelled for ”single use”, the risk analysis should address the hazards associated with reuse as an example of foreseeable misuse.

The results must demonstrate that an appropri- ate risk analysis has been performed and pro- vide a conclusion, with appropriate evidence, that the remaining risks are acceptable when weighed against the intended benefits to the patient. The results of the risk analysis should be reviewed and approved by the manufacturer.

Note: There are a number of published tech- niques for performing a risk analysis. It is recommended that the risk analysis performed in connection with Directive conformity assessment should follow EN 1441.

specification of materials, and manufacturing/special processing
The technical documentation should specify the materials used in the construction of the device, together with the biological safety and biocom-

patibility of materials intended to come into con- tact with the body. Particular attention should be paid where materials are invasive with respect to the body and/or will have prolonged contact with the body.

This may include special processes (e.g. moulding, sterilization) and environmental condi- tions to be used for production (e.g. prevention of particulate contamination or electrostatic dis- charge). In the case of medical devices covered by the IVDD, this should cover characterisation of starting materials.

A general description of the method of manu- facture of the device(s) concerned forms part of the ”Product Description” detailed in 4. This is provided solely to allow a basic understanding necessary for conformity assessment proce- dures.

The majority of procedures, work instructions etc. and the training associated with the proper performing of these tasks forms part of the manufacturer’s quality system.

The technical documentation should specify any ‘special processes’, for example sterilization, the results of which may affect the safety and per- formance of the finished device(s). (See
3.2 (vi)).

specifications, drawings and circuit diagrams for compo- nents, sub-assemblies and the complete product including packaging, where appropriate.
The manufacturer should determine what speci- fications, drawings, diagrams etc. are ap- propriate and sufficient to enable proper manu- facture, installation, maintenance and servicing etc. of the product(s) involved in order to assure the intended characteristics and performances are achieved and maintained.

In some cases, the manufacturer will need de- tailed engineering scale drawings for their prod- uct(s). Whereas such drawings may be neces- sary, for example, for electro-medical devices, it is often sufficient to produce a schematic dia-

gram of, for example, product configuration or kit contents with dimensions and other characteris- tics indicated as appropriate.

Equally, it may be necessary to have drawings for certain components or sub-assemblies but not for others.

the specifications of the checks, tests and trials that are intended to be carried out as part of routine production
The procedures, work instructions etc. relating to the conduct of such checks, tests and trials form part of the manufacturer’s quality system.

the performances and com- patibilities intended by the manufacturer
The manufacturer is required by the directives to identify the characteristics, performances and compatibilities needed to assure the safe and correct operation of the device. A relevant char- acteristic might be, for example, sterility assur- ance of a catheter. A relevant performance might be, for example, the ability of the protective packaging to maintain sterility of that catheter when subject to the stresses associated with transport and storage.

In the case of IVDs, the indication of perform- ances should include those required in connec- tion with analytical performance, for example, to do with sensitivity, specificity, limit of detection, and ratio of false to true results, where relevant.

labelling, including any in- structions for use
The manufacturer is required by the Directives to include in the technical documentation the label, and where appropriate, the instructions for use, together with any changes to these during the lifetime of the product. This should include the information to be given, both by text and the use of symbols, in the final version of the labelling.

The labelling documentation should make clear where particular information will be provided, for example on the device itself or its component parts, on the packaging for each unit, on the sales packaging, or on the leaflet or user manual supplied with one or more devices.

Information may be provided, for example, by means of electronic display screens or synthe- sised voice messages.

Note: Several European standards provide use- ful guidance for devices covered by the MDD, AIMD and IVDD on the circumstances in which particular information is required, and the type of information needed. Additional information may be required by applicable regulations (for exam- ple, the dangerous substances Directive), or product-specific standards, or to indicate the presence of natural latex, materials of animal origin etc.

Note: See also NB-MED Recommendation NB- MED/2.5.2/Rec3 ”Translation procedure”

identification of ‘shelf-life’ re- flected by any ‘use by’ date, or other ‘lifetime’ of the device(s)
In certain cases, such restrictions on use will re- flect a time-related deterioration in characteris- tics that are important to product safety and performance. In other cases, however, the re- strictions will be based on other considerations. The ‘lifetime’ of an active device, for example, may be determined by the period for which the manufacturer will support the device by way of availability of spare parts, manuals, training, service/repairs etc.

Note: See also the NB-MED Recommendation NB-MED/2.2/Rec3 ” ‘Use-by’ date for medical devices”. Guidance on the verification of the sta- bility of IVDs is provided in a European standard (in draft).

Results of Bench Testing	Bench testing includes in-vitro/animal studies,
simulated use testing and validation of software and the results of special processes (e.g. sterili- zation validation report(s)).

Note: Testing should follow a pre-defined proto- col, which should include the parameters to be measured, measuring and test equipment to be used including calibration arrangements, statisti- cal treatment of results and acceptance criteria,

together with necessary formal approval of the report

Clinical data	Clinical data includes data from market experi- ence of the same or similar devices (particularly relevant to ‘well established’ devices), prospec- tive clinical investigations and information from the scientific literature.

The scientific literature will often relate to medi- cal devices other than those being assessed.
The manufacturer must therefore establish the extent to which the scientific literature is relevant to his device(s). The results from bench testing may be used to establish the extent to which the characteristics of the device(s) being assessed are similar to those of the device(s) covered by the scientific literature, and therefore the rele- vance of that scientific literature.

The manufacturer should make clear where clinical data is being used to demonstrate con- formity with each of the applicable essential re- quirements for the particular device(s) con- cerned.

Note: Guidance on the ”Evaluation of Clinical Data” is provided in the NB-MED Recommenda- tion, NB-MED/2.7/Rec3.

Note: In the case of IVDs, investigations in a clinical environment are described as ”perform- ance evaluation studies”.

Documentation and reporting of Design Changes
The technical documentation should include re- cords of each design change and the reasons for these, together with any associated verifica- tion/validation data. The documentation should include evidence for believing that the change achieves the desired effect, and that the device continues to comply with the requirements of the Directive.

Where the technical documentation or part thereof has been submitted to the Notified Body

in connection with conformity assessment in- volving design - or type - examination, the manufacturer is required to inform the Notified Body of substantial changes and obtain further approval.

Note: Guidance on the ”Reporting of design changes and changes of the quality system” is provided in the NB-MED Recommendation, NB- MED/2.5.2/Rec2. This Recommendation de- scribes the system to be applied by the manu- facturer for classifying changes as ‘substantial’ and so must be approved by the Notified Body, and what must be included as a supplement to the technical documentation already submitted.

# Administrative Details

Subject	Guidance

Declaration of Conformity	Upon completion of all other steps required for
conformity assessment, the manufacturer is ex- pected to provide a written declaration that the device(s) concerned meet the provisions of the Directive which apply to them, regardless of whether or not a Notified Body is involved in the conformity assessment.

Note: The declaration must be the final step in the relevant conformity assessment procedure. It would not be possible for example, for the manufacturer to issue a final declaration under MDD Annex II until the Notified Body had ap- proved the quality system, and additionally for class III devices approved the design. It may be helpful however for the manufacturer to prepare a draft declaration of conformity for Notified Body review.

The declaration should make clear under what Directive(s) and Annex(es) it is made, and the product(s) to which it relates. It should also pre- sent the name and address of the manufacturer, and in the case of devices for which the manu- facturer is not resident in the Community, addi- tionally the name and address of the authorised representative of the manufacturer established within the Community.

Note: See also NB-MED Consensus Statement S/01/99 on ”Declaration of conformity”

Application for Conformity Assessment
Where the manufacturer lodges an application for approval of the device type it should include the information referred to in 4.1 (i), (ii), (iii), (v),
(vi) and (vii), together with the relevant parts of the technical documentation.

Note: Particular Notified Bodies may have de- veloped standard forms to be used for applica- tion for conformity assessment. Where available,

these should be used.

Declaration that no other Noti- fied Body is used in Conformity Assessment
Where conformity assessment involves a Noti- fied Body within the application for conformity assessment, the manufacturer is expected to provide a written declaration that no application has been lodged with any other Notified Body for the same product type and conformity assess- ment route.

Notified Body Decisions and Reports
The manufacturer is required by the Directive to keep the Notified Body decisions and reports at the disposal of the national authorities for a pe- riod at least five years after last product to which they relate has been manufactured. These deci- sions and reports include the Notified Body cer- tificates of approval of the quality system and the device design/type, together with any supple- mental approvals of ‘substantial’ changes to the quality system and device design/type.

Note: For guidance on the classification of changes as ‘substantial’ and their reporting to the Notified Body see NB-MED Recommenda- tion NB-MED/2.5.2/Rec2.

Manufacturer’s undertaking on procedure to review post-pro- duction experience
The manufacturer is required by the Directives to institute and keep up to date a systematic pro- cedure to review experience gained from devices in the post-production phase and to implement appropriate means to apply any necessary cor- rective action and to notify the Competent authorities of relevant incidents.

Note: See MedDev 2.12/1 ”Guidelines on a medical device vigilance system”.

# Recommended structure of the technical documentation

To enable effective management of the technical documentation for placing on the market or market surveillance purposes, excessive paperwork should be avoided. To achieve this and to facilitate the task of manufacturers, Notified Bodies and Compe- tent Authorities, it is proposed that the technical documentation is subdivided into two parts.

# Part A of the technical documentation

The first part (A) would consist of a summary of the essential technical data relevant to the conformity assessment procedures, including in particular the data listed be- low.

Note: Such a ‘Summary’ would contain details of compliance with regulations, stan- dards etc. and act as an overview and ‘road map’ to the relevant parts of the manu- facturer’s technical documentation and their status and location. It may also include or refer to appropriate certifications, quality system procedures, reports, declarations etc.

Note: Where the ‘technical documentation’ is to be submitted to the Notified Body to enable assessment of conformity of the device(s) with the requirements of the rele- vant Directive, all that is needed is a statement of the manufacturer’s Quality Policy and an overview of the quality system (e.g. quality manual). Clearly, it is neither ap- propriate nor feasible to include copies of all the various Standard Operating Proce- dures which may apply to the company’s manufacturing and quality assurance ar- rangements, since these documents are constantly being revised and may be exten- sive.

The ‘part A’ of the technical documentation may serve as the basis for submission to Competent Authorities and for class I enforcement etc., and also as the starting point for answering enquiries where there are problems etc.

Subject	Guidance

the name and address of the ‘manufacturer’ within the meaning of the Directive(s).
Where the manufacturer is not resident in the Community, additionally the name and address of the authorised representative of the manu- facturer established within the Community

identification of the device(s) covered by the ‘summary documentation’
This should include the trade or proprietary name(s) if applicable, the common or usual name(s), the device classification(s) and rule(s)

assigned by the manufacturer in accordance with the relevant Directive Annex.

the name(s) and address(es) of the facilities
This should include all the facilities involved in the design and manufacture of the particular de- vice(s) covered by the ‘Summary Statement’

the name and address of any Notified Body involved

a statement of the conformity assessment procedure being followed

the declaration of conformity	This would include the manufacturer’s Declara-
tion of Conformity with the essential require- ments of the medical devices Directives

Note: Where a Notified Body has been involved in the conformity assessment procedure, the No- tified Body’s certificate(s) relating to the product design/type and/or the manufacturer’s quality system should be included

a brief description of the de- vice(s)
The description should include the intended pur- pose(s) and indications for use, together with a listing of any accessories

label and instructions of use

a statement of relevant regu- lations
This should make clear the regulations which the products comply with, together with reference to any third party certifications and approvals

identification of technical standards with which compli- ance is claimed
This should include reference to any third party certifications

a brief statement of the bench testing performed and clinical data obtained
This should make clear how the results of bench testing and clinical data are used to demonstrate compliance with the Directive(s), and make ref- erence to relevant part(s) of the manufacturer’s technical documentation

# Part B of the technical documentation

The second part (B) would consist of remaining technical documentation detailing the risk analysis, the test reports, information concerning the quality manual, plans, descriptions of the products and processes, standards applied, etc. as detailed in the previous section.

If the manufacturer does not follow this recommended two-part structure of the tech- nical documentation and unless the details given in the declaration of conformity or in the certificate of conformity appear sufficient, the authorities could request sub- mission of the complete technical documentation or part thereof according to the re- quirements for relevant purposes.

The manufacturer should be able to demonstrate where and how the various parts of the technical documentation are held and maintained. The technical documentation may be held as decentralised separate documents or together in a single location. Equally, it may be maintained in electronic or hardcopy form.

# Availability of the technical documentation

The technical documentation should be kept at the disposal of the Competent Authorities for inspection and control purposes.

This obligation is incumbent upon the ”manufacturer” within the meaning of the medical devices Directives who places the product on the market under his own name.

Any person responsible for placing a product on the market but not in possession of the technical documentation should be capable of:

stating where the technical documentation is situated;

presenting the technical documentation as soon as possible on request from the Competent Authorities.

However, the name and address of the person in possession of the technical docu- mentation need not be expressly mentioned on the product or on its packaging.

The technical documentation should not be requested systematically by the Notified Body or Competent Authority. In general, it should be requested only during checks made for placing on the market or market surveillance purposes by the Competent Authorities.

To facilitate these tasks part A of the technical documentation should be capable of ready transmission to the relevant authorities. This will eliminate or reduce the need to refer to the remaining documentation identified in part B.

The complete technical documentation has to be able to be collated but this action should only required where total assessment of conformity of the device with the Di- rectives is necessary.

If the Competent Authorities request the technical documentation, the first part of the technical documentation should be submitted immediately, allowing a reasonable time for transmission. Additional time will be required for the submission of the rele- vant sections of the second part of the documentation, taking into account its volume and form (written, computerized, ...) and location.

This organisation of the technical documentation should enable avoidance of re- peated submission of the same technical documentation by the same manufacturer to different authorities.

The technical documentation should be kept for at least five years from the last date of manufacture of the product. The manufacturer should, however, also consider

product liability issues when determining the time and content of technical docu- mentation maintained for obsolete devices.

# Language of the technical documentation

The technical documentation should be maintained in the language selected by the manufacturer either by himself or in agreement with a Notified Body where applica- ble. The Competent Authority may request presentation of the first part of the techni- cal documentation in the official language of the Competent Authority but should not do so if the authority can understand the documentation or its contents in another language. Where a translation is required, the person in possession of the docu- mentation should be allowed extra time to submit the first part of the documentation to the authorities. Part B of the technical documentation should be accepted in the language established by the manufacturer.

Note: See also NB-MED Recommendation NB-MED/2.5.2/Rec3 ”Translation procedure”